SOFTWARE / FIRMWARE
REQUIREMENTS AND DESIGN DOCUMENT

BLE Link and App–Tracker Communication Protocol

Project / Product: NOARK rehabilitation gaming platform — app ↔ tracker link
Document ID: 0.1.0
Version: 0.1
Date: 2026-10-22
Status: Draft

Purpose: Define how the NOARK Godot app and the camera tracker communicate over Bluetooth Low Energy (BLE): transport, connection handling, message protocol (v3), configuration, verification and risks. UDP is covered only where it differs. Sections of the template that do not apply to the communication link are omitted (7, 11, 14, 15, 21); numbering follows the template. Unknowns are marked TBD.

# Table of Contents

1. General Information

1.1 Document Revision Table

1.2 Acronyms and Abbreviations

1.3 Review and Approval

2. Introduction

3. Objectives

4. Scope

5. Deliverables

6. Requirements

6.1 Functional Requirements

6.2 Non-Functional Requirements

6.3 Hardware and Platform Requirements

6.4 Software and Development Environment

8. Standards and References

9. Materials, Tools and Test Equipment

10. System Architecture

10.4 Transport and Connection

10.5 Communication Protocol and Messages

12. Software / Firmware Design

13. Modes of Operation

16. Resource Management

17. Configuration and Customization

18. Testing and Validation

19. Other Considerations and Development Plan

20. Safety and Security Considerations

22. Associated Risks and Control Measures

23. Acronyms, BOM Codes and Appendices

# 1. General Information

| Project / Product Name | NOARK — rehabilitation gaming platform for stroke patients |
| --- | --- |
| Software / Firmware Component | App ↔ tracker communication link: BLE GATT transport (primary), UDP (alternate) and the shared application protocol. Peripheral side: Python tracker service (noark-tracker) on the tracking board. Central side: Godot 4.5 app (Android tablet, or the board itself over UDP). |
| Document Title and ID | BLE Link and App–Tracker Communication Protocol / ID TBD |
| Document Version / Status | 0.1 / Draft |
| Date of Issue / Last Review | 2026-10-22 / TBD |
| Prepared By / Reviewed By | TBD / TBD |
| Development Owner | Internal team (NOARK, biorehabilitation group) |
| Target Release / Milestone | Current: protocol v3 (HELLO version byte 3), branch trunk_demo. Release milestone TBD. |
| Related Documents | NA |

Additional details / notes:

This document describes the implementation as it stands in the repository on 2026-10-22. Where code and docs/PROTOCOL.md disagree, the code wins (see 19, Known Limitations).

# 1.1 Document Revision Table

| Revision | 0.1 |
| --- | --- |
| Date | 2026-10-22 |
| Author | TBD |
| Changes Made | Initial draft, written from PROTOCOL.md and the link source (protocol v3). |

# 1.2 Acronyms and Abbreviations

| Acronym | See table below |
| --- | --- |
| Meaning / Definition | See table below |

Additional details / notes:

| Acronym | Meaning |
| --- | --- |
| ACK | Acknowledgement message (0x82) answering one app command |
| APK | Android application package (the tablet build of the Godot app) |
| ATT | Attribute Protocol — BLE layer that carries GATT reads, writes and notifications |
| BLE | Bluetooth Low Energy |
| BlueZ | Linux Bluetooth stack used by the tracker (through bless / dbus-next) |
| Central / Peripheral | BLE roles: the central (app) scans and connects; the peripheral (tracker) advertises and serves GATT |
| GATT | Generic Attribute Profile — services and characteristics |
| HCI | Host Controller Interface — commands sent straight to the Bluetooth controller |
| MTU | Maximum Transmission Unit (ATT MTU; 247 bytes requested by the app) |
| seq | 16-bit command sequence number used to match ACKs and drop duplicates |
| UDP | User Datagram Protocol — alternate transport when the app runs on the board |
| UUID | 128-bit universally unique identifier of a GATT service or characteristic |

# 1.3 Review and Approval

| Prepared By / Date | TBD |
| --- | --- |
| Technical Review / Date | TBD |
| Quality / Safety Review (if applicable) | TBD |
| Approved By / Date | TBD |

# 2. Introduction

| Background | NOARK patients play rehabilitation mini-games by moving a hand-held device carrying AprilTag markers. A tracking board (Radxa Dragon Q6A, Qualcomm SoC, two cameras) runs the tracker service, which detects the markers, solves the device pose and the patient's trunk posture, and sends them to the game app. The app normally runs on an Android tablet and reaches the board over BLE; when the app runs on the board itself it uses UDP on localhost. Both use the same message protocol. |
| --- | --- |
| Problem / Need | The games need the hand position at camera frame rate with low latency, a way to control the tracker (session, tracking on/off, trunk neutral, settings, recordings, restart) and a link that recovers by itself from drops and tracker restarts, without wires between tablet and board. |
| Intended Users / Operators | Developers and maintainers of the app and tracker; test engineers. Therapists and patients use the link only indirectly (connection status dot, Restart tracker button). |
| Document Purpose | Record the BLE transport, connection behaviour and message protocol so the two sides can be maintained, verified and changed together, and so risks of the link are visible. |

# 3. Objectives

| Primary Objective | Deliver every solved hand pose from the tracker to the app over BLE, and every app command to the tracker exactly once, with automatic recovery from link loss. |
| --- | --- |
| Secondary Objectives | One protocol for BLE and UDP; the transport only moves bytes. Tracker state visible in the app (cameras off / starting / tracking / failed). Cameras and tracking off unless a game or tracking screen is open (power, heat). Messages larger than one BLE notification (snapshot image, storage list) carried transparently. |
| Success Measures | Position messages received at the tracker's frame rate (hand tracking measured ~25–31 fps) — app-side rate measurement TBD. End-to-end latency: target TBD, not yet measured. Reconnect after a tracker restart or out-of-range event without user action (≈3 s detection + ≤10 s scan window + connect). No command applied twice; every command answered or reported as failed within 3 s. |

# 4. Scope

| 4.1 In Scope | BLE GATT service, advertising (incl. raw HCI fallback), scanning and connection handling. Application protocol v3: framing, commands, ACKs, heartbeats, chunking, message formats. Link layer on both sides (pyscripts/link.py, Main_screen/Scripts/tracker_link.gd). Link-related settings (stream_type, ble_device_name, udp_port/udp_host). UDP transport, only where it differs from BLE. |
| --- | --- |
| 4.2 Out of Scope | Marker detection, pose solving, calibration and trunk algorithms (what the messages contain is in scope, how it is computed is not). Game logic, session CSV logging, server upload (sender_raspberryPI.py). Android plugin internals beyond the behaviour the link relies on. |
| Dependencies | BlueZ on the board, accessed through bless / dbus-next; hcitool for the advertising fallback. GdAndroidBLE Android plugin (submodule external/GdAndroidBLE) on Android; gdble GDExtension on desktop. systemd unit noark-tracker; sudoers entries noark-ble and noark-tracker. |
| Assumptions / Constraints | One app connected to a tracker at a time. App and tracker built from the same code revision (same protocol VERSION). Board Bluetooth controller AIC8800: kernel rejects BlueZ advertising data (see 10.4). No pairing or encryption is used today (see 20). |

# 5. Deliverables

| Source Code | github.com/SujithChristopher/NOARKGames (branch trunk_demo); Python 3.13 (tracker) and GDScript (app); Kotlin Android plugin in external/GdAndroidBLE. |
| --- | --- |
| Executable / Firmware Build | Android APK exported from Godot (GUI export, arm64-v8a); Linux ARM64 export for the board; tracker runs from source as the noark-tracker systemd service. |
| Technical Documentation | PROTOCOL.md, this document, CLAUDE.md sections on the link. |
| Test Deliverables | Manual test tools (see 18); test plan, results and defect list TBD. |
| Release Documentation | TBD |
| Knowledge Transfer | TBD |
| Ownership / Licensing | IP ownership TBD. Third-party: bless, bleak, dbus-next (Python); gdble, GdAndroidBLE (Godot). Licence list TBD. |

# 6. Requirements

| Requirement ID | FR-xxx (functional), NFR-xxx (non-functional) — full list in the tables below. |
| --- | --- |
| Requirement Description | One testable requirement per row of the tables below. |
| Rationale / Source | PROTOCOL.md and the implementation (protocol v3); real-time game input. |
| Priority | All FR: Must. NFR-001/002/004: Should, targets TBD. |
| Input / Trigger | App start, link loss, user actions that send commands, camera frames. |
| Expected Behaviour / Output | See each row. |
| Acceptance Criteria | See each row. |
| Verification Method | Test / inspection / demonstration (see 18 and Appendix C). |
| Owner / Status | Owner TBD; status per row. |

Additional details / notes:

Functional requirements

| ID | Requirement | Acceptance criteria | Verif. | Status |
| --- | --- | --- | --- | --- |
| FR-001 | Tracker advertises as a connectable BLE peripheral named ble_device_name with service 4e4f4152-4b00-0000-0000-000000000000 whenever no app is connected. | Visible to a scanner within one 10 s scan; visible again ≤2 s after a disconnect. | Test | Implemented |
| FR-002 | App scans in 10 s windows until it finds the device by name, connects, discovers the service, subscribes to …4b01 and writes to …4b02. | App reaches CONNECTED without user action after start or after a drop. | Test | Implemented |
| FR-003 | On Android the app requests ATT MTU 247 before service discovery. | Notifications up to 244 bytes arrive whole. | Test | Implemented |
| FR-004 | First command is HELLO with the protocol version; a version mismatch is refused with a reason and the app retries every 5 s. | Mismatched APK logs 'tracker speaks protocol v3' and does not stream. | Test | Implemented |
| FR-005 | App commands are sent one at a time and re-sent every 0.3 s until ACKed (10 tries; HELLO unlimited). The tracker ACKs every command; a repeat of the last seq gets the same ACK without being applied again. | No duplicate effects under forced ACK loss; failure reported after 3 s. | Test | Implemented |
| FR-006 | App sends HEARTBEAT every 0.1 s; tracker sends TRACKER_HEARTBEAT when idle for 0.5 s and on any state change. 3 s of silence ends the link on either side. | Both sides detect loss within 3 s ± one frame. | Test | Implemented |
| FR-007 | After a reconnect the app sends HELLO and then the open SESSION again. | Session id at tracker restored without user action. | Test | Implemented |
| FR-008 | While tracking, the tracker sends one POSITION (39 bytes) per frame with a solved pose; none when no marker is visible. | POSITION count at app = solved frames at tracker. | Test | Implemented; rate check TBD |
| FR-009 | Messages larger than the transport limit (244 B BLE, 1400 B UDP) are split into CHUNKs (≤255) and reassembled before the app sees them. | Snapshot and storage list arrive intact. | Test | Implemented |
| FR-010 | Cameras run only after CONFIG track=on and switch off camera_idle_off_s (20 s) after track=off or link loss. | TRACKER_HEARTBEAT state follows idle → starting → tracking → idle. | Test | Implemented |
| FR-011 | RESTART is ACKed, then the tracker exits (code 75) and systemd restarts it. | App reconnects by itself after the restart. | Demo | Implemented |
| FR-012 | On app exit the app sends STOP and waits up to 1 s for its ACK. | Tracker closes the session and keeps running. | Test | Implemented |
| FR-013 | settings.json stream_type selects BLE or UDP on both sides; same protocol on both. | Both values work end to end. | Test | Implemented |

Non-functional requirements

| ID | Requirement | Acceptance criteria | Verif. | Status |
| --- | --- | --- | --- | --- |
| NFR-001 | End-to-end latency camera frame → app position. | Target TBD | Test | Not measured |
| NFR-002 | Sustained position rate over BLE equal to the tracking rate (~25–31 fps) while TRUNK_STATUS (~11 Hz) is also sent. | ≤1% loss over 10 min (proposed) | Test | Not measured |
| NFR-003 | App works on Android 12+ tablets with runtime Bluetooth permissions. | Fresh install connects after granting permissions. | Test | Implemented |
| NFR-004 | Link confidentiality / authentication. | TBD (see 20) | Analysis | Open |
| NFR-005 | Protocol change is made in both wire-format files together with a VERSION bump. | Code review checklist | Inspection | Process |

# 6.1 Functional Requirements

| Function / Feature | App ↔ tracker link: connect, hand-shake, command delivery, data streaming, recovery. |
| --- | --- |
| Normal Behaviour | Tracker advertises; app scans, connects, subscribes; HELLO → ACK; SESSION → ACK. App sends CONFIG track=on when a game opens; tracker streams POSITION / TRUNK_STATUS; heartbeats both ways keep the link alive. |
| Boundary Conditions | Largest single BLE message 244 bytes; larger ones chunked, at most 255 chunks (61,200 bytes). seq is u16 and wraps; chunk msg_id is u8 and wraps. Tracker queues at most 64 unread writes (oldest dropped beyond that). Device name compared exactly; advertised name truncated to 29 bytes in the HCI fallback. |
| Failure Behaviour | No ACK after 10 tries (3 s): command reported failed ('no answer'), next command proceeds. ACK 'no HELLO': tracker restarted — app treats link as lost and starts over with HELLO. 3 s silence: link lost; queued commands reported failed ('link lost'); app rescans / re-HELLOs. Unknown or malformed messages are ignored and logged by the tracker. |
| Related Interface / Data | GATT service and characteristics (10.4), messages (10.5), settings.json (17). |
| Acceptance Test | Run 18.2 / 18.5 scenarios: connect from cold, play a game, restart tracker, walk out of range and back — app resumes streaming each time without user action. |

# 6.2 Non-Functional Requirements

| Performance | POSITION: 39 bytes per solved frame (~25–31 fps ⇒ ~1.2 kB/s). TRUNK_STATUS: CSV, ~11 Hz when trunk tracking is on. App heartbeat 10 Hz (1 byte); tracker heartbeat at least every 0.5 s. Latency target TBD (not measured). |
| --- | --- |
| Reliability / Availability | Tracker is a systemd service, restarted on failure (Restart=on-failure, 5 s). Advertising continues after every disconnect; app rescans indefinitely. Commands reliable via app-level ACK/retry; streamed data (POSITION, TRUNK_STATUS) best effort. |
| Security / Privacy | No pairing, bonding or encryption today; characteristics are open. Patient id and a person-outline JPEG travel in clear. See 20. |
| Power Consumption | Board: cameras off when idle (~98% CPU free, 71 °C) vs tracking (~170% CPU, 85–87 °C). BLE connection interval left to OS defaults; tablet battery impact TBD. |
| Memory / Storage | Tracker write queue 64 messages; app keeps partial chunked messages per msg_id until complete. The link stores nothing on disk. |
| Compatibility | Android: plugin minSdk 24, compileSdk 36; runtime permissions for Android 12+. Godot 4.5; gdble GDExtension on Windows / Linux x86_64 / macOS. Board: Linux (kernel 7.0.11-qcom), BlueZ, Python ≥3.13. App and tracker must speak the same protocol VERSION (3). |
| Maintainability | Wire format in exactly one file per side (protocol.py / link_protocol.gd); transport chosen in one place per side; transports interchangeable (send / receive / reset). |
| Usability / Accessibility | Main menu status dot shows link and tracker state; Restart tracker button; no BLE setup steps for the user beyond granting permissions. |

# 6.3 Hardware and Platform Requirements

| Target Hardware / Device | Tracker: Radxa Dragon Q6A (Qualcomm SoC) with AIC8800 Wi-Fi/Bluetooth and two CSI cameras. App: Android tablet (model TBD) with BLE. |
| --- | --- |
| Electrical / Interface Limits | N/A (wireless link only). |
| Operating Environment | Clinic / home therapy room; tablet within BLE range of the board (range TBD, typically a few metres). RF environment with other 2.4 GHz devices. |
| Host / Mobile Device Requirements | Android with BLE and permissions BLUETOOTH_SCAN, BLUETOOTH_CONNECT, ACCESS_FINE_LOCATION granted; support for ATT MTU 247 (falls back to the default otherwise). |
| Hardware Dependencies | Board Bluetooth controller reachable as hci0; extended-advertising capable controller (HCI LE Set Extended Advertising commands) for the fallback advertiser. |

# 6.4 Software and Development Environment

| Language / Framework | Python 3.13 (tracker); GDScript, Godot 4.5 (app); Kotlin (Android plugin); Rust (gdble GDExtension). |
| --- | --- |
| IDE / Compiler / SDK | Godot 4.5 editor; uv for Python; Gradle + Java 17 + Android SDK for the plugin (./gradlew assemble in external/GdAndroidBLE). |
| Operating System / Runtime | Board: Linux ARM64, systemd, BlueZ. Tablet: Android. |
| Libraries / Dependencies | bless ≥0.3.0 (GATT peripheral on BlueZ), dbus-next — tracker bleak ≥3.0.1, simplepyble ≥0.12.1 — test centrals GdAndroidBLE (submodule) — Android central gdble GDExtension (addons/gdble) — desktop central Licences TBD |
| Build Configuration | Android: GUI export recommended (headless export may not load GDExtensions). Tracker: run from source by the noark-tracker unit. |
| Version Control / Branching | Git, GitHub; main + feature branches; protocol changes reviewed together on both sides. |
| Configuration Management | {DOCUMENTS}/NOARK_demo/settings.json on each device (see 17); sudoers files on the board. |

# 8. Standards and References

| Applicable Standards / Regulations | TBD — to be confirmed (e.g. IEC 62304 / IEC 81001-5-1 if the product is classed as medical device software). |
| --- | --- |
| Product / Hardware Specifications | Radxa Dragon Q6A documentation; AIC8800 module documentation (TBD). |
| Interface / Protocol Documents | docs/PROTOCOL.md — NOARK app ↔ tracker protocol (VERSION 3) Bluetooth Core Specification (GATT, ATT, HCI LE extended advertising) |
| External Libraries / SDK Documentation | bless, BlueZ D-Bus API, Android BluetoothGatt API, Godot 4.5 docs. |
| Internal Procedures | TBD |

# 9. Materials, Tools and Test Equipment

| Development Tools | Godot 4.5 editor, uv, Android Studio / Gradle, SSH to the board. |
| --- | --- |
| Test Equipment | Board with cameras; Android tablet; second BLE phone with a generic BLE scanner app (e.g. nRF Connect) for advertising / GATT checks. |
| Hardware Samples | Board and tablet revisions TBD. |
| Test Software / Utilities | pyscripts/ble_test.py — tracker-like peripheral with a fake hand circle, no cameras pyscripts/tests/test_ble_receiver.py, test_btplug_connect.py — minimal centrals (HELLO, print, STOP) journalctl -u noark-tracker — tracker log ([BLE], [LINK] lines) btmon / hcitool on the board — HCI trace |
| Access Required | Repository access; SSH to the board (sudo for systemctl restart). No secrets recorded here. |

# 10. System Architecture

| 10.1 Component Structure | Three layers on each side: wire format (encode/decode), link (reliability, liveness, chunking) and transport (BLE or UDP, bytes only). Transports share one interface (send / receive / reset / max_message on Python; send / status / stop / received on Godot). Map below. |
| --- | --- |
| 10.2 Main Workflow / Flowchart | See Appendix A sequence. Tracker runs permanently; the app connects, hand-shakes, opens a session, turns tracking on and off as screens change, and stops on exit. |
| 10.2.1 Startup and Initialization | Tracker: systemd starts tracker.py → reads settings.json → BLEStreamer registers the GATT service (bless/BlueZ) → advertising via BlueZ, or RawAdvertiser over HCI if BlueZ advertising fails → ready within 10 s or start fails. Cameras stay off. App: GlobalScript._start_link() picks the transport from stream_type → Android: request BLUETOOTH_SCAN / BLUETOOTH_CONNECT / ACCESS_FINE_LOCATION, wait 1 s, initialise adapter → scan. |
| 10.2.2 Main Operational Flow | Scan → match name → connect → (Android) MTU 247 → discover services → subscribe …4b01 → CONNECTED. HELLO(3) → ACK 'v3' → SESSION(patient) if logged in → CONFIG / TRUNK / STORAGE as the user acts. While track=on: POSITION per frame, TRUNK_STATUS, occasional SNAPSHOT / STORAGE_LIST (chunked). Heartbeats: app every 0.1 s; tracker after 0.5 s without other traffic. |
| 10.2.3 State Machine | App transport: IDLE → SCANNING → CONNECTING → CONNECTED (subscribed, writable). App link: not connected → (HELLO ACKed) connected → (3 s silence, transport down, 'no HELLO') not connected. Tracker camera state (TRACKER_HEARTBEAT): 0 idle, 1 starting, 2 tracking, 3 failed (retry in 5 s). Details in 13. |
| 10.3 Data Flow | Tracker → app: POSITION, TRUNK_STATUS, SNAPSHOT, STORAGE_LIST, ACK, TRACKER_HEARTBEAT. App → tracker: HELLO, SESSION, STOP, CONFIG, TRUNK, STORAGE, RESTART, HEARTBEAT. Position path below. |
| 10.4 Interfaces and Connections | Wireless: BLE GATT service 4e4f4152-4b00-… (10.4 Transport). Local: settings.json on each side; systemd unit noark-tracker; sudoers noark-ble (hcitool) and noark-tracker (systemctl restart). |
| 10.5 Error and Recovery Flow | Command lost → retry 0.3 s ×10 → reported failed. ACK lost → app retries → tracker re-sends cached ACK (same seq), not re-applied. Tracker restarted → next command ACKed 'no HELLO' (or 3 s silence) → app re-HELLOs and re-sends SESSION. BLE drop → app rescans (10 s windows); tracker resets link state, keeps advertising (raw HCI re-enabled every 2 s while unconnected). Write fails with a closed handle (Android) → treated as disconnect. |

Additional details / notes:

Component map

| Layer | Tracker (Python) | App (Godot) |
| --- | --- | --- |
| Wire format | pyscripts/protocol.py | Main_screen/Scripts/link_protocol.gd |
| Link (ACKs, retries, chunks, liveness) | pyscripts/link.py | Main_screen/Scripts/tracker_link.gd |
| BLE transport | pyscripts/ble_streamer.py (BLEStreamer, RawAdvertiser) | Main_screen/Scripts/ble_transport.gd (GdAndroidBLE / gdble) |
| UDP transport | pyscripts/udp_streamer.py | Main_screen/Scripts/udp_transport.gd |
| Transport choice | tracker.py _init_transport() | global_script.gd _start_link() |
| Consumers | tracker.py command handler (_apply_setting, TRUNK, STORAGE, RESTART) | GlobalScript (positions), TrunkMonitor (trunk), Settings / Storage screens |

Data flow (hand position)

cameras → tag detection → pose solve (tracker.py)

→ protocol.encode_position()  (39 bytes)

→ Link.send()  → BLEStreamer.send()  → notify on …4b01 (bless / BlueZ)

~~ BLE ~~

→ GdAndroidBLE characteristic_notified → BleTransport.received

→ TrackerLink._on_received → message(kind, msg)

→ GlobalScript._apply_position → net_x / net_y / net_z → games, raw CSV

# 10.4 Transport and Connection

| Connection Type | BLE (GATT over ATT), primary. UDP over localhost (127.0.0.1:8000) when the app runs on the board — same protocol. |
| --- | --- |
| Peripheral / Central or Client / Server Roles | Tracker = GATT peripheral / server (bless on BlueZ). App = GATT central / client (GdAndroidBLE on Android, gdble BluetoothManager on desktop). |
| Connection Parameters | Service / characteristic UUIDs: table below. Device name: ble_device_name (default NOARK_Tracker), must match on both sides. ATT MTU 247 requested by the Android plugin on connect ⇒ 244-byte message limit; device_connected is raised only after the MTU exchange, so nothing races it. Connection interval / latency / supervision timeout: OS defaults, not set (TBD). |
| Discovery / Pairing / Authentication | Discovery: app scans 10 s at a time and connects to the first device whose advertised name equals ble_device_name; rescans until found. Pairing / bonding / encryption: none. Authorisation is only the protocol HELLO version check. |
| Timeout / Retry Policy | Scan window 10 s, repeated. BLE server start ≤10 s or the tracker fails to start. Command retry every 0.3 s, 10 tries (3 s); HELLO retried indefinitely; refused HELLO → wait 5 s. Liveness: 3 s without any message ⇒ link lost (both sides). Chunks paced 15 ms apart (BlueZ drops a notification whose value is replaced before sending). |
| Disconnect / Reconnect Behaviour | App: on disconnect or connection failure clears state and rescans; on reconnect sends HELLO then the current SESSION again. Tracker: on 3 s silence ends the session (cameras off after camera_idle_off_s), clears the write queue, forgets the last ACK and keeps advertising. App exit: STOP with ≤1 s wait for ACK (or one unacknowledged STOP on hard quit); tracker keeps running. |
| Power / Throughput Requirements | Steady load while tracking: POSITION ~1.2 kB/s + TRUNK_STATUS (~11 Hz CSV) + heartbeats. Chunked messages: 240 payload bytes per 15 ms ⇒ ~16 kB/s max (a 5 kB snapshot ≈ 0.33 s). Power targets TBD. |
| Current vs Target Implementation | Current: open link, name-based discovery, OS-default connection parameters, raw HCI advertising on the AIC8800. Target (proposed, TBD): LE Secure Connections pairing + encrypted characteristics; connect by bonded address; tuned connection interval; measured latency budget. |

Additional details / notes:

GATT database (tracker = peripheral)

| Item | UUID | Properties | Use |
| --- | --- | --- | --- |
| Service | 4e4f4152-4b00-0000-0000-000000000000 | Primary | NOARK link ('NOARK' = 4e 4f 41 52 4b) |
| To app | 4e4f4152-4b01-0000-0000-000000000000 | Read, Notify | One notification = one tracker → app message. Initial value 0x00 (ignored). |
| To tracker | 4e4f4152-4b02-0000-0000-000000000000 | Write, Write Without Response | One write = one app → tracker message. App uses write-without-response when offered. |

Advertising

| Parameter | Value |
| --- | --- |
| Primary path | BlueZ LE advertising through bless (name = ble_device_name, service UUID) |
| Fallback (RawAdvertiser) | Used when BlueZ advertising fails (AIC8800: kernel rejects MGMT Add Extended Advertising Data → Invalid Parameters). GATT service stays BlueZ's. |
| HCI commands | LE Set Extended Advertising Parameters (0x0036), Data (0x0037), Scan Response (0x0038), Enable (0x0039) via 'sudo -n hcitool -i hci0 cmd 0x08 …' |
| Advertising set | Handle 5 (clear of BlueZ's sets from 1); legacy ADV_IND, connectable + scannable |
| Interval / channels / PHY | 100 ms (0x00A0 × 0.625 ms); channels 37/38/39; 1M PHY; public address |
| Advertising data | Flags 0x06 + Complete Local Name (≤29 bytes) |
| Scan response | Complete list of 128-bit service UUIDs (the service UUID) |
| Re-advertise | Every 2 s while no central is connected (a connection ends the set) |
| Permission | /etc/sudoers.d/noark-ble: radxa may run /usr/bin/hcitool only |

# 10.5 Communication Protocol and Messages

| Protocol Overview | Binary, message-oriented protocol (VERSION 3) shared by BLE and UDP. One type byte per message; app commands carry a sequence number and are acknowledged; heartbeats both ways; large messages chunked. Spec: docs/PROTOCOL.md. |
| --- | --- |
| Message / Packet Structure | See framing and tables below. |
| Command Table | App → tracker table below (HELLO, SESSION, STOP, CONFIG, TRUNK, STORAGE, RESTART; HEARTBEAT unacknowledged). |
| Response / ACK / NACK | Every command gets one ACK: ok=1 (applied) or ok=0 (refused) with a text reason. No separate NACK type. Reasons include: 'tracker speaks protocol v3' (version), 'no HELLO' (tracker restarted), per-command validation messages. App reports 'no answer' after 10 tries and 'link lost' for queued commands on link loss. |
| Timing / Ordering | One command in flight; queue is FIFO. This also keeps Android from refusing overlapping BLE writes. Retry 0.3 s; heartbeat 0.1 s (app) / 0.5 s idle (tracker); silence 3 s. Tracker queues up to 64 writes in arrival order and handles them each loop iteration. Streamed data has no sequence number; latest value wins. |
| Version Negotiation | HELLO carries the app's VERSION. Equal → ACK ok 'v3'; different → ACK not ok 'tracker speaks protocol v3', app logs it and retries every 5 s. No downgrade: both sides must be built from the same revision (rebuild the APK after a protocol change). |
| Example Messages | Appendix B. |
| Logging / Traceability | Tracker: stdout to the systemd journal — [BLE] (advertising, connect, streaming), [LINK] (HELLO, version mismatch, unknown/malformed message), [NET]. App: Godot console — [BLE], [LINK] lines, push_error / push_warning for refusals. No persistent message log or capture tool (TBD). |
| Compatibility | Unknown message types: ignored and logged by the tracker; the app forwards unrecognised types to listeners, which ignore them. Messages shorter than their header are ignored. Incompatible changes require a VERSION bump in protocol.py and link_protocol.gd together. |

Additional details / notes:

Framing

–  Every message: 1 type byte + payload. Little-endian. Text is UTF-8, no terminator. Type 0x00 = nothing (ignored).

–  App commands (all except HEARTBEAT): type · seq u16 · body. The tracker answers each with an ACK carrying that seq.

–  ACK: 0x82 · seq u16 · ok u8 (1/0) · detail (UTF-8).

–  CHUNK: 0x85 · msg_id u8 · index u8 · total u8 · part. Part size = max_message − 4 (240 bytes on BLE).

App → tracker

| Type | Name | Body | ACK detail |
| --- | --- | --- | --- |
| 0x01 | HEARTBEAT | (no seq, no body); every 0.1 s | — |
| 0x02 | HELLO | version u8 (3) | 'v3', or why refused |
| 0x03 | SESSION | patient id | the id |
| 0x04 | STOP | — | — |
| 0x05 | CONFIG | key=value | key=value as applied, or why refused |
| 0x06 | TRUNK | neutral ¦ snapshot ¦ select=<n> | the command, or why refused |
| 0x07 | STORAGE | list ¦ delete=<path> | 'listing' / 'deleting <path>'; result later as STORAGE_LIST |
| 0x08 | RESTART | — | 'restarting'; then exit 75 and systemd restart |

Tracker → app

| Type | Name | Payload |
| --- | --- | --- |
| 0x80 | TRACKER_HEARTBEAT | state u8: 0 idle, 1 cameras starting, 2 tracking, 3 cameras failed. Sent after 0.5 s without other traffic and at once on a change. |
| 0x81 | POSITION | <BH9f, 39 bytes: ref_id u16, centroid f32×3 (m, camera frame), rvec f32×3 (Rodrigues), tvec f32×3 |
| 0x82 | ACK | seq u16 · ok u8 · detail |
| 0x83 | TRUNK_STATUS | CSV: state, level, lvl_flex, lvl_lat, lvl_axi, flex, lat, axi, progress, has_neutral, reason, capture_result, people, locked, q_how, q_rms_mm, q_frac, raw_pts, mask_px, held, hz, method, doubt, reference (neutral = angles from captured neutral, model = from the model's own frame) |
| 0x84 | SNAPSHOT | u16 json_len · json {w, h, locked, people:[[[x,y],…],…]} · JPEG (~160 px) |
| 0x85 | CHUNK | msg_id u8 · index u8 · total u8 · part |
| 0x86 | STORAGE_LIST | JSON {total, free, event, recordings:[{path, patient, session, name, bytes, active}]} |

POSITION layout

| Offset | Field | Type | Notes |
| --- | --- | --- | --- |
| 0 | type | u8 | 0x81 |
| 1 | ref_id | u16 | reference tag / rigid-body id |
| 3 | centroid x, y, z | f32 ×3 | tracked point, metres, camera frame |
| 15 | rvec | f32 ×3 | Rodrigues rotation of the rigid body |
| 27 | tvec | f32 ×3 | translation of the rigid body, metres |

CONFIG keys (tracker.py _apply_setting)

| Key | Values | Meaning |
| --- | --- | --- |
| track | on ¦ off | App wants tracking now (cameras on). Not stored in settings.json. |
| solver | joint ¦ ransac | Pose solver |
| refine | crop ¦ contour ¦ subpix ¦ none | Corner refinement |
| camera | both ¦ cam0 ¦ cam1 | Cameras used |
| deadband | pixels | Corner deadband |
| trunk | on ¦ off | Trunk tracking |
| trunk_method | icp ¦ msam_r18 ¦ msam_v3rot | Trunk method (switch clears neutral) |
| trunk_reference | neutral ¦ model | Trunk angles from captured neutral or the model's own frame |
| recording | on ¦ off | Camera recording (from the next session) |
| recording_fps / _scale / _chunk_frames | number | Recording parameters (from the next session) |

At each HELLO the tracker re-applies the same keys from its own settings.json (tracker_solver, tracker_refine, tracker_camera, corner_deadband_px, trunk_enabled, trunk_method, trunk_reference, recording*). Other settings need 'sudo systemctl restart noark-tracker'.

# 12. Software / Firmware Design

| 12.1 File / Folder Structure | pyscripts/{protocol, link, ble_streamer, udp_streamer, tracker}.py; Main_screen/Scripts/{link_protocol, tracker_link, ble_transport, udp_transport, global_script}.gd; external/GdAndroidBLE (Android plugin); addons/gdble (desktop GDExtension). |
| --- | --- |
| 12.2 Core Modules | Wire format, link, transport on each side (10.1). |
| 12.2.1 Communication Module | BLEStreamer: bless GATT server on its own asyncio thread; send() schedules the notification on that loop (dbus is not thread-safe); writes go to a 64-entry deque read by receive(). RawAdvertiser: HCI fallback advertiser (10.4). BleTransport (Godot): scan / connect / subscribe / write, Android or gdble backend. |
| 12.2.2 Data Processing Module | protocol.py / link_protocol.gd: encode / decode of every message, struct formats, chunk splitting. |
| 12.2.3 Control / Application Logic | link.py: dispatches commands once, caches the last ACK, enforces HELLO first, heartbeat, chunk pacing. tracker_link.gd: command queue, retry, liveness, reconnect + session resume, chunk reassembly, status for the UI. |
| 12.2.4 Storage / Logging Module | Link logs only (13 / 10.5). STORAGE command delegates to pyscripts/storage.py on a worker thread. |
| 12.2.5 UI / API Module (if applicable) | Signals from TrackerLink: message(kind, msg), command_done(kind, ok, detail), connected_changed, tracker_state_changed. API: send_command(), set_session(), stop_session(), close(). |
| 12.2.6 Diagnostics / Update Module | Main-menu status dot (link + tracker state); Restart tracker (RESTART, or sudo systemctl restart when the tracker does not answer and the app is on the board). |
| Module Interfaces | Python transport: send(bytes), receive() → bytes, reset(), stop(), max_message. Godot transport: send(PackedByteArray), status(), stop(), signal received(msg). |
| Concurrency / Scheduling | Tracker: main loop thread (camera, link.poll / tick); BLE asyncio loop thread; chunk sender daemon thread per large message. App: Godot main thread, _process each frame; plugin callbacks arrive as signals; link nodes use PROCESS_MODE_ALWAYS so they run while the game is paused or quitting. |
| Memory / Resource Management | Bounded write queue (64); chunk buffers freed when a message completes or index 0 of the same msg_id restarts it. |

# 13. Modes of Operation

| Mode Name / ID | Tracker camera state, reported in TRACKER_HEARTBEAT: 0 Idle, 1 Starting, 2 Tracking, 3 Failed. |
| --- | --- |
| Purpose | Keep cameras and processing off unless the app needs tracking (CPU, heat), and show the user what the tracker is doing. |
| Entry Conditions | Driven by CONFIG track=on/off. The app sends track=on while a scene under res://Games/ is open or a node holds tracking (GlobalScript.hold_tracking), else track=off. Every HELLO starts idle. |
| Configuration Parameters | camera_idle_off_s (20 s) in settings.json. |
| Behaviour / Sequence | Table below. |
| Inputs / Outputs | In: CONFIG track. Out: TRACKER_HEARTBEAT state (at once on change), POSITION while tracking. |
| Metrics / Data Recorded | Tracker journal: [FPS] every 5 s while tracking; recordings cover tracked stretches only. |
| Transitions / Exit Conditions | See table. |
| Fault / Safe Behaviour | Camera failure → Failed, retried every 5 s; link loss → session closed and cameras off after the idle delay. |

Additional details / notes:

| State (TRACKER_HEARTBEAT) | Entry | Behaviour | Exit |
| --- | --- | --- | --- |
| 0 Idle | Service start; track=off + camera_idle_off_s (20 s); link lost | Cameras off; link answers; heartbeats | CONFIG track=on → Starting |
| 1 Starting | track=on while idle | Cameras opening on own thread (2–5 s); link keeps answering | Cameras up → Tracking; error → Failed |
| 2 Tracking | Cameras up | POSITION per solved frame; TRUNK_STATUS if trunk on | track=off → (20 s) Idle |
| 3 Failed | Cameras would not start | Retried after 5 s | Retry → Starting |

# 16. Resource Management

| 16.1 Data Storage Rules | The link stores nothing. STORAGE lists / deletes the board's camera recordings only (<patient>/<session>/Video/rec_* and …/MovementData/recordings/rec_*), never game data, sessions.csv or the recording in progress. |
| --- | --- |
| 16.2 Synchronization / Transfer Rules | N/A for the link. Server sync / upload is a separate process (sender_raspberryPI.py). |
| 16.3 Resource Limits | Message: 244 B (BLE) / 1400 B (UDP); chunked message ≤255 parts (≤61,200 B on BLE). Tracker write queue: 64 messages. seq: u16. chunk msg_id: u8. |
| 16.4 Power Management | Cameras off when not tracking (idle ~98% CPU free, 71 °C vs ~170% CPU, 85–87 °C tracking). BLE power on board / tablet not tuned or measured (TBD). |
| 16.5 Data Integrity and Recovery | Integrity relies on BLE link-layer CRC / ATT and UDP checksums; no application checksum. Commands: ACK + retry + seq de-duplication. Streams: best effort; a lost POSITION is replaced by the next. Chunks: a message with a missing part is never delivered; the buffer is reset when index 0 of the same msg_id arrives. |

# 17. Configuration and Customization

| Configuration File / Interface | {DOCUMENTS}/NOARK_demo/settings.json on each device (JSON). Edited from the app's Settings screen (Tracker tab) or by hand on the board. |
| --- | --- |
| Configurable Parameters | Link keys in the table below; tracking keys in 10.5 (CONFIG). |
| Validation / Defaults | Missing keys take the defaults above; an unknown stream_type falls back to ble with a warning on both sides. The device name must match on both sides or the app never connects. |
| Configuration Update | Live keys go to the tracker as CONFIG and are re-applied from its own file at each HELLO; transport keys need 'sudo systemctl restart noark-tracker' on the board. |
| Example Configuration | See below (link keys only). |
| Version Migration | Settings file is seeded from defaults on first run; new keys default when absent. |

Additional details / notes:

| Key | Type | Default | Read by | Change applies |
| --- | --- | --- | --- | --- |
| stream_type | ble ¦ udp | ble | App, tracker | App: live (link restarted). Tracker: service restart |
| ble_device_name | string | NOARK_Tracker | App, tracker | App: next link start. Tracker: restart |
| udp_port | int | 8000 | App, tracker | Restart |
| udp_host | string | 127.0.0.1 | App | Next link start |
| camera_idle_off_s | seconds | 20 | Tracker | Restart |

{

"stream_type": "ble",

"ble_device_name": "NOARK_Tracker",

"udp_port": 8000

}

# 18. Testing and Validation

| 18.1 Unit Testing | None for the link or protocol today (pytest tests cover tracking only; no Godot test framework). Proposed: round-trip tests for protocol.py (encode / decode / chunk) and a fake transport test for link.py (dedup, no-HELLO, heartbeat). |
| --- | --- |
| 18.2 Integration Testing | App ↔ ble_test.py (fake peripheral, no cameras); test_ble_receiver.py / test_btplug_connect.py centrals ↔ real tracker; tablet APK ↔ board. |
| 18.3 System / Functional Testing | Full session on tablet: connect, log in, play each game, trunk neutral / snapshot / select, Settings live changes, Storage list / delete, exit. |
| 18.4 Performance / Endurance Testing | TBD: position rate at app vs tracker [FPS]; latency; 1 h session without drop; chunk throughput. |
| 18.5 Fault / Recovery Testing | Restart tracker mid-game; walk out of range and back; Bluetooth off / on on tablet; kill app; mismatched APK version; deny permissions; board reboot. |
| 18.6 User / UI Testing | Status dot shows scanning / connecting / connected and cameras off / starting / tracking; Restart tracker button. |
| Test Environment | Radxa Dragon Q6A, kernel 7.0.11-qcom, AIC8800; Android tablet model TBD; app and tracker from the same commit. |
| Defect Management | TBD (GitHub issues proposed). |
| Release Acceptance | TBD — proposed: all FR tests pass on the target tablet + board; no open high-severity link defects. |

# 19. Other Considerations and Development Plan

| Implementation Phases | Done: protocol v3 over BLE and UDP, raw HCI advertising fallback, tracking on demand, storage, restart. Next (proposed): link security (20), latency / rate measurement, automated protocol tests. |
| --- | --- |
| Priorities | 1. Security of the open link. 2. Measured performance targets. 3. Automated tests. |
| Migration / Backward Compatibility | No backward compatibility across VERSION: app and tracker are updated together; an older APK is refused at HELLO and must be rebuilt. |
| Known Limitations | No pairing or encryption; any BLE central in range can connect and send commands. App connects by device name only; a second device with the same name could be chosen. Designed for one app at a time; a second central is not refused. BlueZ advertising fails on the AIC8800; fallback needs sudo hcitool and is board-specific. gdble has no Linux ARM64 build, so the app on the board uses UDP. docs/PROTOCOL.md title says v2 while VERSION is 3; GdAndroidBLE plugin/demo still uses the old wire format. Connection parameters not tuned; latency not measured. |
| Maintenance Plan | Owner TBD. Keep docs/PROTOCOL.md and this document in step with protocol changes. |

# 20. Safety and Security Considerations

| Potential Safety Hazards | The device only senses; there are no actuators, so link faults cannot move anything. Indirect: frozen / stale position or a missed trunk COMPENSATING warning during exercise after a link drop; data logged under the wrong patient if SESSION is wrong. |
| --- | --- |
| Safe State / Emergency Behaviour | Link loss detected within 3 s; status dot shows it; tracker closes the session and switches cameras off after 20 s. Game behaviour on link loss: verify per game (TBD). |
| Cybersecurity Threats | Unauthenticated access: any central can read positions / trunk / snapshot and send CONFIG, STORAGE delete=<recording> or RESTART. Spoofing: a device advertising NOARK_Tracker could feed false positions. Eavesdropping: traffic is not encrypted. Denial of service: repeated RESTART or connection hogging. |
| Risk Controls | Current: STORAGE limited to recording paths, never the active one; sudoers entries limited to one command each; version check at HELLO. Proposed: LE Secure Connections bonding with encrypted / authenticated characteristic permissions; connect to bonded address; allowlist on the tracker. |
| Privacy / Sensitive Data | SESSION carries the patient id; SNAPSHOT carries a small JPEG with person outlines; both in clear over BLE. Encrypting the link is recommended. |
| Applicable Compliance Evidence | TBD |

# 22. Associated Risks and Control Measures

| Risk ID / Description | R-01 — open, unauthenticated BLE link (more risks in the table below). |
| --- | --- |
| Potential Effect | Patient data exposure; commands from an unknown device (delete recording, restart tracker). |
| Likelihood / Severity | TBD (project risk matrix). |
| Risk Level | TBD (project risk matrix). |
| Control / Mitigation | Proposed: LE Secure Connections bonding, encrypted characteristics, bonded-address connect. |
| Verification of Control | Attempt connection / write from an unbonded device (TBD). |
| Owner / Residual Risk | TBD. |

Additional details / notes:

| ID | Risk | Effect | Control / mitigation | Verification |
| --- | --- | --- | --- | --- |
| R-01 | Open, unauthenticated GATT link | Data exposure; unwanted commands | Proposed bonding + encryption | Security test TBD |
| R-02 | Spoofed tracker by name | False positions to game | Proposed bonded address | Test TBD |
| R-03 | Kernel / BlueZ update changes advertising | App cannot find tracker | Raw HCI fallback; re-test after OS update | 18.5 |
| R-04 | App and tracker on different VERSION | No connection | HELLO refusal with reason; rebuild APK | 18.5 |
| R-05 | Notification loss under burst | Chunked message incomplete | 15 ms chunk pacing; user may retry (snapshot / list) | 18.4 |
| R-06 | Link drop mid-trial | Gap in position data | 3 s detection, auto reconnect, session resend | 18.5 |

# 23. Acronyms, BOM Codes and Appendices

| Acronym / Code | See 1.2. |
| --- | --- |
| Description | Appendices below. |
| Appendix A: Diagrams | Connection sequence below; component map and data flow in 10. |
| Appendix B: Message Examples | Below. |
| Appendix C: Test Matrix | Below. |
| Appendix D: Release / Version History | Below. |

Additional details / notes:

Appendix A: Sequence

app (central)                         tracker (peripheral)

|  scan 10 s, match ble_device_name     |  advertising

|-- connect, MTU 247, discover, subscribe ->|

|-- HELLO(seq, 3) ---------------------->|  re-apply live settings

|<----------------------- ACK(seq,1,'v3')|

|-- SESSION(seq, patient) -------------->|  open logs / recordings

|<------------------------- ACK(seq,1,id)|

|-- CONFIG(seq, 'track=on') ------------>|  cameras start (2-5 s)

|<---------------- TRACKER_HEARTBEAT(1/2)|

|<----- POSITION, TRUNK_STATUS ... ------|  per frame

|-- HEARTBEAT every 0.1 s -------------->|

|-- STOP(seq) -------------------------->|  close session, keep running

|<------------------------- ACK(seq,1,'')|

Appendix B: Message examples (hex)

| Message | Bytes |
| --- | --- |
| HEARTBEAT (app) | 01 |
| HELLO, seq 0x1234, version 3 | 02 34 12 03 |
| ACK ok 'v3' | 82 34 12 01 76 33 |
| ACK refused 'no HELLO' (tracker restarted) | 82 35 12 00 6E 6F 20 48 45 4C 4C 4F |
| CONFIG 'track=on', seq 0x1236 | 05 36 12 74 72 61 63 6B 3D 6F 6E |
| TRACKER_HEARTBEAT tracking | 80 02 |
| POSITION ref 0, centroid (0.10, -0.05, 0.45), rvec 0, tvec (0.08, -0.02, 0.42) | 81 00 00 CD CC CC 3D CD CC 4C BD 66 66 E6 3E 00 00 00 00 00 00 00 00 00 00 00 00 0A D7 A3 3D 0A D7 A3 BC 3D 0A D7 3E |
| CHUNK msg 7, part 0 of 5 | 85 07 00 05 + 240 bytes |

Appendix C: Test matrix

| Requirement | Test |
| --- | --- |
| FR-001, FR-002, FR-003 | 18.2 connect from cold with scanner app and tablet |
| FR-004 | 18.5 mismatched APK |
| FR-005, FR-007, FR-011 | 18.5 restart tracker mid-game; unit test link.py (proposed) |
| FR-006 | 18.5 out of range / Bluetooth off |
| FR-008, NFR-001, NFR-002 | 18.4 rate and latency (TBD) |
| FR-009 | 18.3 trunk snapshot, storage list |
| FR-010 | 18.6 status dot through menu → game → menu |
| FR-012, FR-013 | 18.3 exit app; switch stream_type |

Appendix D: Version history

| Protocol | Notes |
| --- | --- |
| v3 (current) | HELLO version byte 3; STORAGE, RESTART, TRACKER_HEARTBEAT state. Earlier versions: see git history of docs/PROTOCOL.md. |